## Slide 1
mPEMMR (SA)
User Guide

## Slide 2
System Requirements
Installation
Program Execution
Troubleshooting

## Slide 3
I. System Requirement
RAM 4GB or More
JDK 1.5 or More (64bit Recommended)
TPP (http://tools.proteomecenter.org/wiki/index.php?title=Software:TPP)
msconvert (Installed as option when TPP is installed)
Thermo Foundation 2.0 (for Q-Exactive)
Xcalibrabur 2.2 (for Q-Exactive)

## Slide 4
II. Installation
Unzip mPEMMR_1.1.8.zip
  log4j (directory)
  LIB (directory)
  RAPID1.07 (directory)
  mPEMMR_1.1.8.cmd (Command Executable File)
  mPEMMR_param.txt (mPEMMR Parameter File)
  UserGuide.pptx (User Guide)
mPEMMR_1.1.8.cmd : Open file in text editor to modify as below
  @echo off
  SET PEMMR_HOME=D:\Programs\mPEMMR_1.1.8 (mPEMMR.cmd Change to file existing Directory Address)
  SET CLASSPATH=%PEMMR_HOME%\LIB\mPEMMR_1.1.8.jar;%PEMMR_HOME%\msfile_3rd_party\jmzreader-1.2.0.jar;%PEMMR_HOME%\msfile_3rd_party\mzxml-parser-1.3.2.jar;%PEMMR_HOME%\msfile_3rd_party\xxindex-0.11.jar;%PEMMR_HOME%\msfile_3rd_party\com.springsource.org.apache.commons.logging-1.1.1.jar;%PEMMR_HOME%\msfile_3rd_party\com.springsource.org.apache.commons.io-1.4.0.jar;%PEMMR_HOME%\msfile_3rd_party\log4j-1.2.13.jar;%PEMMR_HOME%\msfile_3rd_party\jcommon-1.0.16.jar;%PEMMR_HOME%\log4j\log4j-1.2.15.jar;
  set MEM=8G
  if ""%1""=="""" goto example
  goto exec1
  :example
  echo Example) mPEMMR_1.1.8 [mPEMMR_Params.txt]
  goto end
  :exec1
  java -Xmx%MEM% pemmr_sa.PEMMRExec %*
  goto end
  :end

## Slide 5
III. Program Execution
Modify mPEMMR_param.txt as below
Run Command line
Change directory file of PEMMR.cmd (In this case, D:\Programs\PEMMR_SA_Release\1.1)
Run the program typing PEMMR.cmd pemmr_sa_param.txt
PEMMR will create xxxx_PEMMR.mgf which will be used for peptide search
mPEMMR_param.txt
#################### External Programs ###################
RAPID=D:/Programs/mPEMMR_1.1.8/RAPID1.07/RAPID.exe
MSCONVERT=C:/Inetpub/tpp-bin/msconvert.exe
MZXML2SEARCH=C:/Inetpub/tpp-bin/MzXML2Search.exe
################## UMC Create ##################
RAW_FILE_URL= D:/Programs/mPEMMR_1.1.8 /Example/N1_5ug_100cm_300min_1_091112.raw
UMC_ALGORITHM=1
SCAN_COUNT=2
HOLE_COUNT=5
MASS_TOLERANCE=10
MIN_MASS=0
MIN_INTENSITY=0
################## Spectrum Filtering ##################
SCAN_RANGE=5
UMC_LINK_TOLERANCE=10
### since ver 1.1.7 (ppm) ###
ISOTOPE_ENVELOPE_TOLERANCE=5
"Multiplexed Post-Experimental Monoisotopic Mass Refinement (mPE-MMR) to increase sensitivity and accuracy in peptide identifications from tandem mass spectra of co-fragmentation" Madar, I. H.; Ko, S.-I.; Kim, H.; Mun, D.-G.; Kim, S.; Smith, R.; Lee, S.-W. Anal. Chem. 2017, 89(2), 1244-1253.

## Slide 6
III. Program Execution (Editing pemmr_sa_param.txt )
mPEMMR_param.txt
#################### External Programs ###################
RAPID=D:/Programs/mPEMMR_1.1.8/RAPID1.07/RAPID.exe
MSCONVERT=C:/Inetpub/tpp-bin/msconvert.exe
MZXML2SEARCH=C:/Inetpub/tpp-bin/MzXML2Search.exe
################## UMC Create ##################
RAW_FILE_URL= D:/Programs/mPEMMR_1.1.8 /Example/N1_5ug_100cm_300min_1_091112.raw
UMC_ALGORITHM=1
SCAN_COUNT=2
HOLE_COUNT=5
MASS_TOLERANCE=10
MIN_MASS=0
MIN_INTENSITY=0
################## Spectrum Filtering ##################
SCAN_RANGE=5
UMC_LINK_TOLERANCE=10
### since ver 1.1.7 (ppm) ###
ISOTOPE_ENVELOPE_TOLERANCE=5
RAPID program is provided with mPEMMR and it is located in PEMMR_HOME sub directory
MSCONVERT and MZXML2SEARCH is installed with TPP
RAW_FILE_URL is URL of RAW file under excution
Caution : All directories typed in mPEMMR_param.txt file should be entered using “/”
e.g.)RAPID=D:\Programs\PEMMR_SA_Release\1.1\RAPID1.07\RAPID.exe (X)
RAPID=D:/Programs/PEMMR_SA_Release/1.1/RAPID1.07/RAPID.exe (O)

## Slide 7
IV. Troubleshooting
Stop running with the message of "java.lang.ClassNotFoundException"
Open PEMMR.cmd file in text editor and make sure the dir of “SET PEMMR_HOME location” entered properly.
Q. Execution is interrupted with the message "Exception in thread "main": java.lang.OutOfMemoryError: Java heap space"
Open PEMMR.cmd in text editor and increase MEM to 8G.
Q. RAPID doesn’t run properly and the file size of resultant files (_isos.csv, _scans.csv) is 0.
A. RAW file has not read properly by the program, Re-Install the Thermo foundation.
If mPEMMR program stopped generating result files, need to delete the existing generated files and run the PEMMR once again